Supplemental Materials

Magnetic Resonance Imaging (MRI) Details

Scan Sequences (Core Protocol)

- 3-Plane / Tri-Planar Scout

- Axial T2-Star / GRE

- Axial T2 FLAIR

- Axial functional connectivity MRI (fcMRI) — eyes open

- Sagittal 3D Accelerated MPRAGE/IRSPGR

- Axial Diffusion Weighted Imaging (DWI)

- Axial T2 FSE/TSE

Optional Sequence: Multiband fcMRI (eyes open)

Acquisition Protocols: https://adni.loni.usc.edu/data-samples/adni-data/neuroimaging/mri/mri-scanner-protocols/

Acquisition Parameters

|  | Axial_T2_STAR | Axial_T2_FLAIR | Sagittal 3D Accelerated MPRAGE | Axial DWI | Axial fcMRI (Eyes Open) | Axial T2 TSE with Fat Sat |
| --- | --- | --- | --- | --- | --- | --- |
| Time | 4:11 | 4:05 | 5:12 | 2:13 | 10:00 | 3:26 |
| Voxel Size | 0.9×0.9×4.0 mm³ | 0.9×0.9×5.0 mm³ | 1.1×1.1×1.2 mm | 1.1×1.1×4.0 | 3.3×3.3×3.3 mm | 0.9×0.9×4.0 mm³ |
| Slices | 44 | 35 | 176 | 40 | 48 | 44 |
| FOV | 220 | 220 | 270 | 220 | 212 | 240 |
| Slice Thickness | 4 | 5 | 1.2 | 4 | 3.3 | 4 |
| TR | 650 | 9000 | 2300 | 10100 | 3000 | 4000 |
| TE | 20 | 91 | 2.95 | 100 | 30 | 78 |
| Avg | 1 | 1 | 1 | 4 | 1 | 1 |
| Flip | 20 | 150 | 9 | n/a | 80 | 150 |
| TI | n/a | 2500 | 900 | n/a | n/a | n/a |
| Matrix | 256x192 | 256x256 | 256x256 | 192x192 | 64x64 | 256x224 |

Acquisition Geometry Requirements: All scans were acquired in either the straight axial or straight sagittal plane, with no oblique angulation permitted under any circumstances. Coverage included the entire brain from the inferior cerebellum to the vertex, and for MPRAGE acquisitions the nose was included in the field of view to prevent image‑folding artifacts.

Scanner Requirement: 3 Tesla (3T) field strength only, with a multi-channel head coil.

QC Measurements Performed: The quality‑control team evaluated phantom scans for linear‑scaling accuracy, gradient nonlinearity, and signal‑to‑noise ratio, which represented the primary performance parameters monitored across sites. The specific pulse‑sequence parameters (TR, TE, voxel size, and related settings) were stored in the electronically transferred protocol files.

Amyloid PET Details

Radiopharmaceutical & Dosing

- Tracer: Florbetapir F 18 (Amyvid™)

- Injected dose: 10 mCi (370 MBq) ± 10%, administered as a single IV bolus within 20 seconds, followed by a saline flush

- Minimum acceptable dose: 5 mCi (below this, no injection should proceed)

- No fasting required prior to injection

Scan Timing

- Uptake period: ~50 minutes post-injection (participant rests quietly)

- Participant voids bladder at ~45 minutes post-injection

- Emission scan starts at 50 minutes ± 5 minutes post-injection

Emission Acquisition

- Frame structure: 4 × 5-minute dynamic acquisitions (total 20 minutes)

- Acceptable fallback (if dynamic not available): 1 × 15-minute static scan

- Scanning direction: Head to foot

- Number of beds: 1 (brain only)

- Entire brain must be within the FOV

Reconstruction Parameters

| Parameter | Value |
| --- | --- |
| Randoms correction | On (Prompts – Delays) |
| Scatter correction | On (Monte Carlo simulation based on CT data model) |
| Decay correction | On |
| Attenuation correction | Measured (low-dose CT or rod source transmission) |
| Reconstruction algorithm | Iterative OSEM — 4 iterations, 16 subsets (or RAMLA, depending on scanner manufacturer) |
| Image matrix size | 128 × 128 |
| Zoom | 2 |
| SFOV | 600 mm |
| Slice thickness | 4 mm (matched to CT slices) |
| Post-reconstruction filter | Gaussian, FWHM 5.0 mm |
| Output units | Quantitative (SUV, Bq/mL, or kBq/cc) |

Attenuation Correction by Scanner Type

- PET-only: Rod source transmission scan, 5–10 minutes, acquired prior to emission

- PET/CT: Standard low-dose, non-diagnostic CT of the head — ~55 eff mAs, ~120 kVp, acquired prior to emission

Tau PET Details

Radiopharmaceutical & Dosing

- Tracer: Flortaucipir F 18 (18F-AV-1451)

- Injected dose: 10 mCi (370 MBq) ± 10%, administered as a single IV bolus followed by 15–20 mL saline flush

- Minimum acceptable dose: 6 mCi (below this, no injection should proceed)

Scan Timing

- Uptake period: 75 minutes ± 5 minutes post-injection before emission scanning begins

- Total scan duration: 30 minutes

Emission Acquisition

- Frame structure: 6 × 5-minute dynamic acquisitions (static 1×15 min was not acceptable)

- Minimum usable frames: 3 out of 6 (scans with >4 motion-affected frames are deemed unanalyzable)

- Field of view: Must include entire brain - full cortex and cerebellum; scans missing any brain anatomy were deemed unanalyzable

- Scanning direction: Head to foot

- Number of beds: 1 (brain only)

Reconstruction Parameters

- Algorithm: Iterative (approximating OSEM - 4 iterations, 16 subsets); LOR-RAMLA / 3D-RAMLA for Philips scanners

- Image matrix size: 128 × 128

- Zoom: 2× (or FOV: 350 mm)

- Slice thickness: 4 mm

- Post-reconstruction filter: Gaussian, FWHM 5.0 mm

- Corrections required: randoms, scatter, decay, and measured attenuation

- Output units: Quantitative (SUV, Bq/mL, or kBq/cc)

Attenuation Correction by Scanner Type

- PET-only: Rod source transmission scan, 5-10 minutes, prior to emission

- PET/CT: Low-dose non-diagnostic CT - approximately 55 eff mAs, ~120 kVp

- PET/MR: Manufacturer-recommended MR AC sequence; parameters (voxel size, TR, TE, slice thickness, flip angle, FOV, acquisition time) were required to be constant across all timepoints.