Supplemental Material

Supplement A. Bland-Altman plots by structure

Supplement B. Version Quality Control metrics

| Modality | SNR v8 | SNR v9 | CNR v8 | CNR v9 |
| --- | --- | --- | --- | --- |
| DWI | 2.61 ± 0.83 | 4.71 ± 1.16 | 3.11e+08 ± 1.35e+08 | 3.78e+08 ± 1.81e+08 |
| FLAIR | 3.00 ± 0.53 | 6.23 ± 1.61 | 2.27e+08 ± 1.05e+08 | 3.87e+08 ± 1.34e+08 |
| T1 | 2.86 ± 0.67 | 6.26 ± 1.21 | 2.34e+08 ± 1.12e+08 | 2.89e+08 ± 1.24e+08 |
| T2 | 3.80 ± 0.48 | 5.44 ± 1.14 | 2.99e+08 ± 1.06e+08 | 3.43e+08 ± 1.52e+08 |
| SNR: signal-to-noise ratio, CNR: contrast-to-noise ratio. Metrics were calculated with RadQy GitHub - viswanath-lab/RadQy: RadQy is a quality assurance and checking tool for quantitative assessment of magnetic resonance imaging (MRI) and computed tomography (CT) data.52 |  |  |  |  |

Supplement C. Linear relationship between cortical region size and ICC

Supplement D. Assessment of software version effects on regional brain volumes

Percent volume differences between pre- (v8) and post- (v9) software upgrade scans are shown for each brain region, computed as the difference between mean post- and pre-upgrade volumes normalized by the pre-upgrade mean. Horizontal lines indicate the mean within-version scan–rescan percent differences estimated from the two repeated acquisitions before (and after) the upgrade, representing expected measurement variability. Post-upgrade volumes showed a consistent positive shift across regions, with pre–post differences exceeding scan–rescan variability in most cases.

Supplement E. Complete table of percentage bias

| ID | Percent pre–post shift (bias) |
| --- | --- |
| total intracranial | 5.842450174 |
| left cerebral white matter | 3.916856988 |
| left cerebral cortex | 11.60279576 |
| left lateral ventricle | 25.12057232 |
| left inferior lateral ventricle | 24.76000846 |
| left cerebellum white matter | -4.768534943 |
| left cerebellum cortex | -3.406905397 |
| left thalamus | -21.45148343 |
| left caudate | 15.6310199 |
| left putamen | -6.931395885 |
| left pallidum | 9.509897477 |
| 3rd ventricle | 12.10681504 |
| 4th ventricle | 9.526375639 |
| brain-stem | -2.459654503 |
| left hippocampus | 29.78870172 |
| left amygdala | 24.84364293 |
| csf | 10.09576704 |
| left accumbens area | 13.11243793 |
| left ventral DC | 8.93472665 |
| right cerebral white matter | 2.211720388 |
| right cerebral cortex | 9.97533424 |
| right lateral ventricle | 12.11714844 |
| right inferior lateral ventricle | 23.74885351 |
| right cerebellum white matter | -22.44156037 |
| right cerebellum cortex | -15.05891671 |
| right thalamus | -8.244922037 |
| right caudate | 8.964530675 |
| right putamen | -4.098571478 |
| right pallidum | -10.18157197 |
| right hippocampus | 19.90632023 |
| right amygdala | 1.254635291 |
| right accumbens area | 13.55206943 |
| right ventral DC | 5.039786848 |
| ctx-lh-bankssts | 17.85839169 |
| ctx-lh-caudalanteriorcingulate | 9.631090579 |
| ctx-lh-caudalmiddlefrontal | 8.723977784 |
| ctx-lh-cuneus | 3.938265211 |
| ctx-lh-entorhinal | 11.38695734 |
| ctx-lh-fusiform | 6.316146996 |
| ctx-lh-inferiorparietal | 12.24776273 |
| ctx-lh-inferiortemporal | 19.16877272 |
| ctx-lh-isthmuscingulate | 4.968026512 |
| ctx-lh-lateraloccipital | -0.186658065 |
| ctx-lh-lateralorbitofrontal | 6.395440357 |
| ctx-lh-lingual | -4.702222646 |
| ctx-lh-medialorbitofrontal | 6.878312867 |
| ctx-lh-middletemporal | 14.74176839 |
| ctx-lh-parahippocampal | 5.657384248 |
| ctx-lh-paracentral | 10.25083422 |
| ctx-lh-parsopercularis | 16.49244426 |
| ctx-lh-parsorbitalis | 16.9696702 |
| ctx-lh-parstriangularis | 19.80937988 |
| ctx-lh-pericalcarine | 0.108097258 |
| ctx-lh-postcentral | 18.50869929 |
| ctx-lh-posteriorcingulate | 7.391402377 |
| ctx-lh-precentral | 15.82898561 |
| ctx-lh-precuneus | 8.297563375 |
| ctx-lh-rostralanteriorcingulate | 25.6358274 |
| ctx-lh-rostralmiddlefrontal | 16.27753878 |
| ctx-lh-superiorfrontal | 14.33412808 |
| ctx-lh-superiorparietal | 2.184458159 |
| ctx-lh-superiortemporal | 16.49753387 |
| ctx-lh-supramarginal | 12.63352504 |
| ctx-lh-frontalpole | 13.2423213 |
| ctx-lh-temporalpole | 12.64443333 |
| ctx-lh-transversetemporal | 12.59644803 |
| ctx-lh-insula | 23.57640094 |
| ctx-rh-bankssts | 10.54415882 |
| ctx-rh-caudalanteriorcingulate | 9.935863863 |
| ctx-rh-caudalmiddlefrontal | 1.13979891 |
| ctx-rh-cuneus | 13.99098909 |
| ctx-rh-entorhinal | 6.324621221 |
| ctx-rh-fusiform | 7.142898318 |
| ctx-rh-inferiorparietal | 4.475499882 |
| ctx-rh-inferiortemporal | 5.634125263 |
| ctx-rh-isthmuscingulate | 4.598725817 |
| ctx-rh-lateraloccipital | 13.44383957 |
| ctx-rh-lateralorbitofrontal | 13.24775954 |
| ctx-rh-lingual | 15.64254241 |
| ctx-rh-medialorbitofrontal | 4.758569615 |
| ctx-rh-middletemporal | 3.196951355 |
| ctx-rh-parahippocampal | 18.47204513 |
| ctx-rh-paracentral | 12.46329043 |
| ctx-rh-parsopercularis | 9.456209597 |
| ctx-rh-parsorbitalis | -0.28781615 |
| ctx-rh-parstriangularis | 3.556827469 |
| ctx-rh-pericalcarine | 12.17939685 |
| ctx-rh-postcentral | 11.69315517 |
| ctx-rh-posteriorcingulate | 7.752901238 |
| ctx-rh-precentral | 17.00404439 |
| ctx-rh-precuneus | 5.285766327 |
| ctx-rh-rostralanteriorcingulate | 9.805804757 |
| ctx-rh-rostralmiddlefrontal | 15.68398874 |
| ctx-rh-superiorfrontal | 15.73284309 |
| ctx-rh-superiorparietal | 9.830472239 |
| ctx-rh-superiortemporal | 9.045250067 |
| ctx-rh-supramarginal | 12.01382488 |
| ctx-rh-frontalpole | 5.125633692 |
| ctx-rh-temporalpole | 1.99084688 |
| ctx-rh-transversetemporal | -2.031886717 |
| ctx-rh-insula | 17.5473756 |

Notes: Median bias = 9.806, Median bias % difference pre-T2 = 4.231, Median bias % difference post-T2 = 2.822.